## Report
Medical Physicist's Radiography Performance Test Summary
Site: |  | Survey date:
Location\Detail: |  | Tested By:
X-ray Unit Type: |  | Service Provider No.:
Vendor-Model: |  | Reviewed/Submitted By:
F Number: |  | Service Provider No.:
SN: |  | Radiation Detector:
Type of Survey: |  | Detector SN:
Calibration date:
|  |  | Evaluation | MDH Rule No.
1. Radiographic Unit Assembly Evaluation |  |  | Pass
2. Collimation Assessment
Source-indicator distance (SID) is accurate within 2% |  |  | Pass | 1100.5A
Deviation between x-ray field and light field is less than 2% of SID. |  |  | Pass | 1100.5B
Deviation between x-ray field and collimator setting is less than 2% of SID. |  |  | Pass | 1100.5D
Deviation between x-ray field center and image receptor is < 2% of SID. |  |  | Pass | 1100.5C
Light field intensity is at least 160 lux. |  |  | Pass | 1100.5N
Positive beam limitation is functioning properly. |  |  | Pass
3. Radiographic generator assessment
Timer accuracy is < 5% |  |  | Pass | 1100.5I
mR/mAs within 10% of baseline |  |  | Pass | 1100.5F
4. kVp Accuracy and Reproducibility
Measured kVp is within ± 5% of nominal kVp. |  |  | Pass | 1100.5E
kVp coefficient of variation is < 0.02 |  |  | Pass
X-ray output coefficient of variation is < 0.02 |  |  | Pass | 1100.5E
Filtration is greater than 2.3 mm Al. |  |  | Pass | 1100.5J
5. mR/mAs Linearity
mR/mAs linearity is less than 0.1. |  |  | Pass | 1100.5G
6. Automatic Exposure Control (AEC) system performance
Phototimer compensation for kVp and patient thickness is adequate. |  |  | Pass | 1100.5O
Minimum optical density is at least 1.0 |  |  | na | 1100.5P&S
Density control function operates within acceptable limits. |  |  | Pass | 1100.5M
Exposure reproducibility is within acceptable limits. |  |  | Pass | 1100.5L
Exposures for left, right and center chambers are set appropriately. |  |  | Pass
Medical Physicist's Recommendations for Quality Improvement:
The machine passed all performance tests conducted.
Radiography Equipment Evaluation
Site: | 0 | Survey date: | 0
Equipment
X-ray Unit Type: | 0 | Radiation Detector: | 0
Vendor-Model: | 0 | Detector SN: | 0
Phantom: | CIRS Model 903 | Calibration date: | 0